## Parameters
Parameter # | Parameter name | Data type | unit | Default value | Limit High | Limit Low | Description
0 | ParamSN0 | ASCII
1 | ParamSN1 | ASCII
2 | ParamSN2 | ASCII
3 | ParamSN3 | ASCII
4 | ParamSN4 | ASCII
5 | ParamSN5 | ASCII
6 | ParamSN6 | ASCII
7 | ParamSN7 | ASCII
8 | ParamButtonDebounceTime | uint16 | ms | 100 |  |  | Debouncer timer for three buttons on joystick
9 | ParamMotorspeedUpdatingPeriod | uint16 | ms | 100 |  |  | How often the joystick voltage is used to update the stepper speed
10 | ParamPZTspeedUpdatingPeriod | uint16 | ms | 100 |  |  | How often the joystick voltage is used to update the PZT speed
11 | ParamMaxCoarseSpeed | uint16 | 1/8 step per second | 1600 |  |  | Maximum coarse speed for stepper, default 1600/8=200 full steps per second
12 | ParamMinCoarseSpeed | uint16 | 1/8 step per second | 16 |  |  | Minimum coarse speed for stepper, default 16/8=2 full steps per second
13 | ParamMaxFineSpeed | uint16 | 1/8 step per second | 200 |  |  | Maximum fine speed for stepper, default 200 1/8 steps per second
14 | ParamMinFineSpeed | uint16 | 1/8 step per second | 2 |  |  | Minimum fine speed for stepper, default 2 1/8 steps per second
15 | ParamMaxJoyVoltage | uint16 | mV | 2000 | 2500 | 1500 | Joystick voltage above this parameter will result maximum speed
16 | ParamMinJoyVoltage | uint16 | mV | 250 | 1000 | 1 | No movement if joystick voltage below this value
17 | ParamMaxPZTStepSize | uint16 | DAC value | 13 |  |  | The maximum step size for PZT to move when joystick voltage is above "ParamMaxJoyVoltage". DAC value of 65535 corresponding to 100um movement.
18 | ParamMaxPZTStepTime | uint16 | ms | 1 |  |  | The step time corresponding to "ParamMaxPZTStepSize"
19 | ParamMinPZTStepSize | uint16 | DAC value | 1 |  |  | The minimum step size for PZT to move when joystick voltage is at ParamMinJoyVoltage. DAC value of 65535 corresponding to 100um movement.
20 | ParamMinPZTStepTime | uint16 | ms | 1 |  |  | The step time corresponding to "ParamMinPZTStepSize"
21 | ParamApproachPZTStepSize | uint16 | DAC value | 2 |  |  | Step size during approaching, DAC value of 65535 corresponding to 100um movement
22 | ParamApproachPZTStepTime | uint16 | ms | 1 |  |  | Step time during approaching corresponding "ParamApproachPZTStepSize"
23 | ParamApproachADCThreshold | uint16 | % | 1 |  |  | threshold of percent of ADC absolute value change from initial value that will stop approaching
24 | ParamApproachVoltage | uint16 | mV | 1000 |  |  | Voltage applied during approach.
25 | ParamEnablePauseAfterThresholdDetected | uint16 | N/A | 0 |  |  | Flag to enable stopping without puncture movement when approaching threshold detected, 1: enable, 0: disable
26 | ParamFixedDistanceMove | uint16 | DAC value | 655 |  |  | Fixed distance puncture move after approaching stopped. default 1um (65535/100). DAC value of 65535 corresponding to 100um movement.
27 | ParamRetractStepSize | uint16 | DAC value | 13 |  |  | Step size during retracting, DAC value of 65535 corresponding to 100um movement.
28 | ParamRetractStepTime | uint16 | ms | 1 |  |  | Step time during retracting
29 | ParamRetractStepNum | uint16 | DAC value | 2520 |  |  | Number of steps during retract, 32768 (50um)/13 = 2520 steps, about total 50um movement
30 | ParamAspiratePulseVoltage | uint16 | mV | 500 |  |  | Pulse voltage applied during aspirating
31 | ParamAspiratePulseEndVoltage | uint16 | mV | 0 |  |  | Voltage applied when pulse ends during aspirating
32 | ParamAspiratePulseDuration | uint16 | ms | 1000 |  |  | Aspirate pulse length, when set to 0, the ParamEjectPulseEndVoltage will be applied Indefinitely
33 | ParamInjectPulseVoltage | uint16 | mV | 5000 |  |  | Pulse voltage applied during injecting
34 | ParamInjectPulseEndVoltage | uint16 | mV | 0 |  |  | Voltage applied when pulse ends during injecting
35 | ParamInjectPulseDuration | uint16 | ms | 500 |  |  | Injecting pulse length, when set to 0, the ParamEjectPulseEndVoltage will be applied Indefinitely
36 | ParamEjectPulseVoltage | uint16 | mV | 2000 |  |  | Pulse voltage applied during ejecting
37 | ParamEjectPulseEndVoltage | uint16 | mV | 0 |  |  | Voltage applied when pulse ends during ejecting
38 | ParamEjectPulseDuration | uint16 | ms | 1000 |  |  | Ejecting pulse length, when set to 0, the ParamEjectPulseEndVoltage will be applied Indefinitely
39 | ParamSensingDuration | uint16 | s | 60 |  |  | Not used anymore
40 | ParamHdrADCAverageNum | uint16 |  | 16 |  |  | number of adc data for averaging (1 to 64)
41 | ParamApproachInitValueAcqTime | uint16 | ms | 500 |  |  | Time to wait for getting the initial approach ADC average current
42 | ParamApproachHoveringNum | uint16 |  | 20 |  |  | number of times the needle back stepping while waiting for average value reaching threshold. If exceeding this number without passing threshold, system goes to fault state
43 | ParamApproachMaxStepNum | uint16 |  | 16384 |  |  | number of max steps the needle will go in the whole approach process, together with "ParamMaxPZTStepSize" will define the total distance (50um, 32768 DAC counts) of moving during approach.
44 | ParamEnablePauseAfterPoking | uint16 |  | 0 |  |  | Enable pause after poking, before the pulse is generated
45 | ParamEnablePauseAfterPulsing | uint16 |  | 1 |  |  | Enable pause after the pulsing, before retracting is happening
| ParamHdrADCLineNoiseAveNum | uint16 |  | 20 |  |  | use 20 = 1200/60 for 60Hz noise, use 24 =1200/50 for 50Hz noise, the sampling rate is 1200.
| ParamHdrADCCountsInitOffset | uint16 |  | 0 |  |  | the offset current in counts when no pipette is connected, a calibration can be done to determine this. Calibration command is in status block
| ParamEnable60HzFilter | uint16 |  | 0 |  |  | Enabling the 60Hz filtering for data and control, default is disabled (if enabled, the data in the ADC data block will be filled with the data after averaging)
| ParamHdrADCAverageMethod | uint16 |  | 0 |  |  | 0: moving average, 1: regular average (average method on top of the ADC data block, the result of this will be used to determine 1% in approach process)
To write
| MSC_PARAMS_BLOCK | 0x0000FFFE
| MSC_COMMAND_BLOCK | 0x000FFFFD
Format is [0] = cmd, [1-2] = data to follow (len), [3-511] = command specific data
| cmd | MSC_SET_PARAMS = 2,
| len | N*3 (N is number of parameter to be changed)
| data | 3 bytes,0: nParam, 1-2: value to write
To save parameters, use MSC_SAVE_PARAMETERS=3, len =1, data[0]=1

## Status
Status # | Status name | Data type | unit | Access | Description
0 | StatFirmWareVersion = 0 | uint16 |  | R | Firmware version of the main board
1 | StatTick | uint16 | ms | R | System ticks in ms
2 | StatChecksum | uint16 |  | R | Checksum of the firmware
3 | StatBootVersion | uint16 |  | R | Firmware version of the bootloader
4 | StatBootChecksum | uint16 |  | R | Checksum of the bootloader
5 | StatState | uint16 |  | R | Current state of the system, please refer to State tab
6 | Stat12V | uint16 | mV | R | 12V power supply reading from ADC
7 | Stat5V | uint16 | mV | R | 5V power supply reading from ADC
8 | StatXSpeed0 | LSB of int32 | 1/8 step per sec | R | Speed of X stepper
9 | StatXSpeed1 | MSB of int32 |  | R
10 | StatYSpeed0 | LSB of int32 | 1/8 step per sec | R | Speed of Y stepper
11 | StatYSpeed1 | MSB of int32 |  | R
12 | StatZSpeed0 | LSB of int32 | 1/8 step per sec | R | Speed of Z stepper
13 | StatZSpeed1 | MSB of int32 |  | R
14 | StatXCounts0 | LSB of uint32 | 1/8 step | R | X distance from home position (backward limit) after homing finishes
15 | StatXCounts1 | MSB of uint32 |  | R
16 | StatYCounts0 | LSB of uint32 | 1/8 step | R | Y distance from home position (backward limit) after homing finishes
17 | StatYCounts1 | MSB of uint32 |  | R
18 | StatZCounts0 | LSB of uint32 | 1/8 step | R | Z distance from home position (backward limit) after homing finishes
19 | StatZCounts1 | MSB of uint32 |  | R
20 | StatHdrADCCounts0 | LSB of int32 | ADC counts | R | current ADC counts from header ADC
21 | StatHdrADCCounts1 | MSB of int32 |  | R
22 | StatHdrADCCountsLineNoiseAve0 | LSB of int32 | ADC counts | R | ADC counts after the line noise averaging, if line noise average enabled (ParamEnable60HzFilter=1), this data will be send to ADC data block
23 | StatHdrADCCountsLineNoiseAve1 | MSB of int32 |  | R
24 | StatHdrADCCountsAverage0 | LSB of int32 | ADC counts | R | current average ADC counts from ADC Data block data (the average method can be determined by ParamHdrADCAverageMethod: 0 moving average, 1 regular average
25 | StatHdrADCCountsAverage1 | MSB of int32 |  | R
26 | StatHdrCurrent0 | LSB of float | current in pA | R | Same meaning as above ADC counts, in pA
27 | StatHdrCurrent1 | MSB of float |  | R
28 | StatHdrCurrentLineNoiseAve0 | LSB of float | current in pA | R | Same meaning as above ADC counts, in pA
29 | StatHdrCurrentLineNoiseAve1 | MSB of float |  | R
30 | StatHdrCurrentAverage0 | LSB of float | current in pA | R | Same meaning as above ADC counts, in pA
31 | StatHdrCurrentAverage1 | MSB of float
32 | StatJoyX | int16 | mV | R | current joystick X axis voltage reading
33 | StatJoyY | int16 | mV | R | current joystick Y axis voltage reading
34 | StatJoyZ | int16 | mV | R | current joystick Z axis voltage reading
35 | StatXCountsValid | uint16 |  | R | Indicating whether the X stepper counts is valid (homing done), 1 is valid
36 | StatYCountsValid | uint16 |  | R | Indicating whether the Y stepper counts is valid (homing done), 1 is valid
37 | StatZCountsValid | uint16 |  | R | Indicating whether the Z stepper counts is valid (homing done), 1 is valid
38 | StatXTimerAAR | uint16 |  | R | Auto reload number used to control the timer pulses for X stepper
39 | StatYTimerAAR | uint16 |  | R | Auto reload number used to control the timer pulses for Y stepper
40 | StatZTimerAAR | uint16 |  | R | Auto reload number used to control the timer pulses for Z stepper
41 | StatPIDACCountsCh1 | uint16 |  | RW | PZT control X DAC value, 0-65535 corresponds to 0-10V
42 | StatPIDACCountsCh2 | uint16 |  | RW | PZT control Y DAC value, 0-65535 corresponds to 0-10V
43 | StatPIDACCountsCh3 | uint16 |  | RW | PZT control Z DAC value, 0-65535 corresponds to 0-10V
44 | StatPISenseCh1 | uint16 | mV | R | PI controller sense feedback voltage for X
45 | StatPISenseCh2 | uint16 | mV | R | PI controller sense feedback voltage for Y
46 | StatPISenseCh3 | uint16 | mV | R | PI controller sense feedback voltage for Z
47 | StatPIOverflowCh1 | uint16 |  | R | PI controller overflow feedback TTL for X (1 for high, 0 for low)
48 | StatPIOverflowCh2 | uint16 |  | R | PI controller overflow feedback TTL for Y (1 for high, 0 for low)
49 | StatPIOverflowCh3 | uint16 |  | R | PI controller overflow feedback TTL for Z (1 for high, 0 for low)
50 | StatPIOnTargetCh1 | uint16 |  | R | PI controller on target feedback TTL for X (1 for high, 0 for low)
51 | StatPIOnTargetCh2 | uint16 |  | R | PI controller on target feedback TTL for Y (1 for high, 0 for low)
52 | StatPIOnTargetCh3 | uint16 |  | R | PI controller on target feedback TTL for Z (1 for high, 0 for low)
53 | StatJoySW1 | uint16 |  | R | Joystick mode switch 1 status (1 for high, 0 for low)
54 | StatJoySW2 | uint16 |  | R | Joystick mode switch 2 status (1 for high, 0 for low)
55 | StatJoySW3 | uint16 |  | R | Joystick mode switch 3 status (1 for high, 0 for low)
56 | StatXLimitF | uint16 |  | R | Indicating if the X axis forward limit switch is tripped (1: stripped, 0: not tripped)
57 | StatXLimitB | uint16 |  | R | Indicating if the X axis backward limit switch is tripped (1: stripped, 0: not tripped)
58 | StatYLimitF | uint16 |  | R | Indicating if the Y axis forward limit switch is tripped (1: stripped, 0: not tripped)
59 | StatYLimitB | uint16 |  | R | Indicating if the Y axis backward limit switch is tripped (1: stripped, 0: not tripped)
60 | StatZLimitF | uint16 |  | R | Indicating if the Z axis forward limit switch is tripped (1: stripped, 0: not tripped)
61 | StatZLimitB | uint16 |  | R | Indicating if the Z axis backward limit switch is tripped (1: stripped, 0: not tripped)
62 | StatXOverride | uint16 |  | RW | Indicating X axis limit switch overriding line status (1: high, 0: low)
63 | StatYOverride | uint16 |  | RW | Indicating Y axis limit switch overriding line status (1: high, 0: low)
64 | StatZOverride | uint16 |  | RW | Indicating Z axis limit switch overriding line status (1: high, 0: low)
65 | StatXEnable | uint16 |  | RW | X axis stepper driver enable line status (1: enabled, 0: disabled)
66 | StatYEnable | uint16 |  | RW | Y axis stepper driver enable line status (1: enabled, 0: disabled)
67 | StatZEnable | uint16 |  | RW | Z axis stepper driver enable line status (1: enabled, 0: disabled)
68 | StatHomeEnable | uint16 |  | W | Enable homing of the steppers for all three axis including PZT (any write value > 0 will trigger homing, homing including PZT)
69 | StatCmdStatus | uint16 |  | RW | Command status of the system, refer to "Command" tab for detail
70 | StatClearHeadCommErr | uint16 |  | W | A write of non-zero value clear the system out of "StateErrHeadComm" state
71 | StatStreamingCtrl | uint16 |  | W | To start/stop the streaming of header data (1: start streaming, 0: stop-streaming, 2: no-action)
72 | StatStreamingStatus | uint16 |  | R | Streaming status of header data (1: streaming, 0: non-streaming)
73 | StatImediateVoltage | int16 | mV | RW | Get/set header board immediate voltage (-10V to 10V)
74 | StatPulseVoltage | int16 | mV | RW | Get/set header board pulse voltage (-10V to 10V)
75 | StatPulseEndVoltage | int16 | mV | RW | Get/set header board end voltage when pulse stops (-10V to 10V)
76 | StatPulseDuration | uint16 | ms | RW | Get/set header board voltage pulse duration
77 | StatStartPulse | uint16 |  | W | A write of non-zero value start a pulse in the header
78 | StatCalibrateHdrADCCountsInitOffset | uint16 |  | W | A write of non-zero will start the calibration of current initial offset, basically the 32 bit integer value StatHdrADCCountsAverage will be applied to ParamHdrADCCountsInitOffset, then all the current data in ADC data block will be offseted by ParamHdrADCCountsInitOffset
| MSC_STATUS_BLOCK | 0x0000FFFF
To write | MSC_COMMAND_BLOCK | 0x000FFFFD
Format is [0] = cmd, [1-2] = data to follow (len), [3-511] = command specific data
| cmd | MSC_SET_STATUSES = 1,
| len | N*3 (N is number of parameter to be changed)
| data | 3 bytes,0: nParam, 1-2: value to write

## ADC Data
ADC data of Header can be accessed by reading a USB block:
Block# | 0xFFFC
Block size | 512 byte
Total number of data : | 64
Size of ADC data | 4 byte
data format | lower 3 bytes are ADC counts, MSB byte representing the data index (circling from 0 to 255)
Size of Z position data | 2 byte | 0-65535: corresponding to 100um to 0 um
unused data | 2 byte

## State
State number | State name
0 | StateStopped | no motion on both steppers and PZT
1 | StateHome | state determining next homing axis
2 | StateHomeX | X axis is doing homing
3 | StateHomeY | Y axis is doing homing
4 | StateHomeZ | Z axis is doing homing
5 | StateCoarseMove | Coarse move state
6 | StateFineMove | Fine move state
7 | StatePZTMove | PZT move state. Can accept commands CmdAspirate, CmdInject, CmdEject to go to StateAspirate, StateInject, StateEject respectively and set up corresponding pulse parameters, then return to StatePZTMove. It also accepts CmdApproach to start the approach process
8 | StateApproach | Approach state
9 | StateEndofApproach | End of approach
10 | StateFixedDistanceMove | A fixed distance PZT z axis movement after the end of approach
11 | StateAspirate | Set up pulse voltage, pulse end voltage and pulse duration related to aspirate, then go back to StatePZTMove
12 | StateAspiratePulse | Not been used
13 | StateInject | Set up pulse voltage, pulse end voltage and pulse duration related to inject, then go back to StatePZTMove
14 | StateInjectPulse | Not been used
15 | StateSensing | Go to this state after pulse is generated, if "ParamEnablePauseAfterPulsing" is set to 1, then it will keep streaming without any movement until "ParamEnablePauseAfterPulsing" is set to 0 or a "retract" command comes in. If "ParamEnablePauseAfterPulsing" is 0 originally, the system will go to retract automatically.
16 | StateRetract | Retract, retract step size, step time and step number are all configurable through parameters
17 | StateEject | Set up pulse voltage, pulse end voltage and pulse duration eject, then go back to StatePZTMove
18 | StateEjectPulse | Not been used
19 | StateErrHeadComm | Error state when there is an error of communicating with header
20 | StateErr | Movement error when 1% threshold couldn't be detected when system goes to maximum approach distance allowed. Or the backstep number passes "ParamApproachHoveringNum" without a detection

## Command
The command user can write to " StatCmdStatus" to have system doing certain actions
command name | command number
CmdNone | 0 | no command
CmdApproach | 1 | approach
CmdAspirate | 2 | aspirate
CmdInject | 3 | inject
CmdSensing | 4 | sensing
CmdEject | 5 | eject
CmdRetract | 6 | retract
CmdStop | 7 | stop any movement, stepper or PZT
CmdHomeX | 0x10 | homing of x
CmdHomeY | 0x20 | homing of y
CmdHomeZ | 0x40 | homing of z
CmdHomeXYZ | 0x70 | do homing x, y, and z